Minutes – RK Hall Feedback / Production HWT Calibration Process

September 10, 2018

1:00 PM to 1:30 PM

Attendees: Jason Moser, Dorian Lewis, Eric Hayes, Yan Zhang, Brian Norman, Clint Van Winkle, Sharayu Ghangrekar, John Vaders, James Galante, Gary Hockenbury

Absent: Finch Troxler, Dan Pratt

Situation:

- Thursday September 6th, 2018

  - Dean Arrives at RK Hall to calibrate Standard WheelTracker

  - During Step 5 of the calibration, Dean receives the following error

  - Sharayu sent him

  - Dean attempts to use his laptop, but battery is dead, he leaves to charge the computer

- Friday September 7th 2018

  - Deans computer will not connect to the PLC receiving the following error

RK Hall Solution:

- Dean was going to ship gauge blocks to RTP, and is now going to ship gauge blocks to Pete

- Dorian is going to prepare a computer to do the calibration

- Eric to ship Computer to Pete

- Pete

  - Install CoDeSys on his laptop

  - Receive

    - Gauge Blocks

    - Computer

  - Visit RK Hall early next week if possible

  - Use laptop to check voltage on bailiffs

  - Use WheelTracker computer to calibrate

- Clint to Call Richard at RK Hall

- Michael Dixon to call his contact at RK Hall

Long-Term Solution:

- Clint to change procedure so that calibration is done at RTP with the computer to be shipped out

- John/Eric Technicians to have their own gauge blocks

- Jason/Brian make CoDeSys available on your respective

  - Service drives

  - NPD drives

- Yan will write possible procedure for connecting to PLC when receiving “Unable to connect to PLC” error