Troxler Model 03440A Main Calibration Report

Gauge Serial number:       26813    Calibration Date:      3/22/2021

Density Standard Count:      1817     Water Standard Count:  622   Bay: 57

The density counts used in this calibration are shown hereunder.  "Pos" refers to

the source position below the surface of the standard, in millimeters.

Magnes     Mag/Al     Alumin

Pos      standard   standard   standard

-----      ------     ------     ------

BS        00779      00520      00357

051        02718      01756      01090

102        02758      01686      00967

152        02229      01265      00676

203        01534      00800      00394

254        00964      00460      00218

305        00572      00255      00117

The density system calibration constants and precision for each source position of

this gauge are listed hereunder. The gauge precision (“Prec”) is the measurement

precision under a set of repeatability conditions of measurement (VIM Sec. 2.21.)

Density precision is evaluated at 2002 kilograms per cubic meter.  The calibration

constant B*1000 is listed in units of cubic meter per kilogram.

Pos        A      B*1000       C      Prec

---      ------   -------   --------  -----

BS       3.037   1.22186   -0.07065    10.5

051       8.943   0.99803    0.06399    5.2

102      10.716   1.07240    0.12315    4.6

152      12.239   1.29135    0.05104    4.6

203      12.412   1.52518    0.01645    5.0

254      13.315   1.85808   -0.01486    5.7

305      11.630   2.08894   -0.01408    6.9

The water system calibration counts, constants, precision (“Prec”), and anticipated

errors for this gauge are listed hereunder. The “SR” value refers to a count made

on the magnesium/polyethylene standard with 1.3 mm. shims beneath the gauge.

“SR Error” is the anticipated error in the water system readings due to surface

texture imperfections and “Exerr” is the anticipated error from the combination

of gauge imprecision and texture imperfections.  The calibration constant F is

listed in units of cubic meter per kilogram, and the precision and errors

are listed in units of kilogram per cubic meter.  The precision is evaluated

at 250.0 kilograms per cubic meter.

Magnesium  Mag/Poly

Standard  Standard   SR                              SR

count      count   count     E     F*1000    Prec  Error   Exerr

---------  --------  -----  -------  -------   ----  ------  -----

19       401      389   0.03055  1.06254    5.1  -18.2   14.6

Troxler Model 03440A Main Calibration Report

Gauge Serial number:       26813    Calibration Date:      3/22/2021

Density Standard Count:      1817     Water Standard Count:  622   Bay: 57

Listed hereunder is a range of projected density standard counts at future dates.

These projected ranges are computed based solely upon the known decay rate of the

density source

Lower Limit of        Upper Limit of

Projected density      Projected density

Date          Standard Count        Standard Count

----------     -----------------     -----------------

4/1/2021           1798                  1834

5/1/2021           1794                  1831

6/1/2021           1791                  1827

7/1/2021           1787                  1824

8/1/2021           1784                  1820

9/1/2021           1780                  1816

10/1/2021           1777                  1813

11/1/2021           1774                  1810

12/1/2021           1770                  1806

1/1/2022           1767                  1803

2/1/2022           1763                  1799

3/1/2022           1760                  1796

4/1/2022           1757                  1792

This gauge was calibrated using the method described in Troxler Report

M-708.  The density standards utilized in this calibration listed

hereunder are traceable to the National Institute of Standards &

Technology (NIST) as documented in Troxler Reports R-703 and R-705.

The uncertainties for these standards are expressed as expanded

uncertainty values at approximately the 95% confidence level using a

coverage factor of k = 2.

Density

Expanded

Standard    Standard    Density     Uncert.     Date of     Date of

Material    Serial No.  (kg/m3)     (kg/m3)   Measurement  Expiration

------------ -----------  --------   --------  -----------  -----------

Magnesium     172        1770.7       7.6    11/11/2019   11/11/2024

Mag/Alum     187        2223.2       10     11/11/2019   11/11/2024

Aluminum     171        2703.3       12     11/11/2019   11/11/2024

To account for the influence of the elemental composition of these blocks on

instrument response (as prescribed in ASTM D7759, Section 16.1.5), the

densities for the metallic block are multiplied by elemental composition

correction factors prior to the calculation of the density calibration

parameters shown on Page 1 of this document.  These

correction factors  are 0.988 for magnesium, 0.974 for magnesium/aluminum,

and 0.964 for aluminum.

Troxler Model 03440A Main Calibration Report

Gauge Serial number:       26813    Calibration Date:      3/22/2021

Density Standard Count:      1817     Water Standard Count:  622   Bay: 57

The measurements used to establish the water mass per unit volume values

for the standards utilized in this calibration are traceable to the National

Institute of Standards & Technology (NIST) as documented in Troxler Report

R-704.  The uncertainty for this standard is expressed as an expanded

uncertainty value at approximately the 95% confidence level using a

coverage factor of k = 2.  Water mass per unit volume is defined as the

mass of water contained in a unit volume of material, and is represented in

the table hereunder as 'M'.

M Expanded

Standard    Standard    M Value     Uncert.     Date of     Date of

Material    Serial No.  (kg/m3)     (kg/m3)   Measurement  Expiration

------------ -----------  --------   --------  -----------  -----------

Mag/Poly       170        577.5       40    11/12/2019   11/12/2024

The maximum estimated measurement uncertainties (EMUs) of this gauge are

listed hereunder.  The uncertainties apply to measurements of one minute in

duration on typical soil, and are expressed as expanded uncertainty values at

approximately the 95% confidence level using a coverage factor of k = 2.

These values relate only to the source rod positions for which this gauge has

been manufactured to operate. EMU values are in kilograms per cubic meter.

For DENSITY Measurements:

Max             Max             Max             Max

Pos   EMU       Pos   EMU       Pos   EMU       Pos   EMU

---   ----      ---   ----      ---   ----      ---   ----

BS     33       51     24       76     22      102     20

127     23      152     26      178     22      203     18

229     18      254     17      279     19      305     21

The maximum EMU for M measurements is 41 kilograms per cubic meter.

This instrument was found to be mechanically sound and electronically stable

prior to and after its calibration.  This calibration was performed in an

environment with a temperature between 14.4 ºC and 25.6 ºC.  All data

listed in this report are applicable to this instrument only.  This

instrument was calibrated by MAP using the Three Block Calibration process.

This process is fit for the intended use of the instrument.

This calibration was performed by MAP at Troxler Electronic Labs

11300 Sanders Dr., Suite 7, Rancho Cordova, CA 95742

Special considerations and limitations of use for this device and its

calibration are described in the Manual of Operation and Instruction

provided with this instrument.

Troxler Model 3440 As Left Calibration Report

Gauge Serial number:       26813    Measurement Date:      3/24/2020

Density Standard Count:      1849     Water Standard Count:  622   Bay: 57

Listed hereunder are the results of a gauge accuracy check for this instrument

for soil and asphalt measurement modes.  The results are listed in terms of

measurement error (as defined in VIM Section 2.16), where the measurement error

is the value read by the instrument minus the value assigned to the standard

upon which the measurement was made.  Measurement errors for density are listed

in kilograms per cubic meter, as are measurement errors for water mass per

unit volume measurements (denoted by M.)  “Pos ” refers to the source

position below the surface of the standard, in millimeters.

Magnes     Mag/Al     Alumin

Pos      standard   standard   standard

-----      ------     ------     ------

BS          -4          9          0

051           4          7          0

102           2          7          2

152          -2         -3          0

203          -4         -4          0

254          -6          5         -1

305          -7         -6         -6

M measurement errors:  On mag. =    0 kg/m3   On mag/poly =    3 kg/m3

All measurements were one minute in duration, except for the standard counts,

which were four minutes in duration.  These measurements were made at the

conclusion of the factory calibration for this gauge.

This instrument was evaluated using the method described in Troxler Report

W-700.  The quantity values and uncertainties of the standards used in this

calibration are listed in the Main Calibration Report for this instrument.

The maximum estimated measurement uncertainties (EMUs) of a gauge of this

variety that has undergone a Troxler factory calibration are listed and

in the Main Calibration Report for this instrument.

This calibration was performed in an environment with a temperature

between 14.4 ºC and 25.6 ºC.

This “as left” calibration was performed  by Miles Platner at

Troxler Electronic Labs, 11300 Sanders Dr., Suite 7, Rancho Cordova, CA 95742.

This report shall not be reproduced except in full, without the written approval

of Troxler Electronic Laboratories, Inc.  This process is fit for the intended

use of the gauge, compliant with ASTM D7759, Sections 9.8 and 13.4.

_______________________________________________     ______________

Approved Signatory (Sign and Print)                      Date

Quality Survey

Dear Customer;

Troxler is conducting a survey in order to better serve our customers. Please complete and return this form to the attention of the Quality Manager.

- What service location did you use for this calibration?

California     Florida     Illinois    Louisiana   North Carolina     Arlington, Texas    Houston, Texas

- How would you rate your experience with Troxler? Please circle your answers.

A = Excellent, B = Good, C = Needs Improvement, D = Poor, N/A = Not Applicable

| Is the calibration certificate accurate and error free? | A | B | C | D | N/A |
| --- | --- | --- | --- | --- | --- |
| How satisfied are you with the information on and the format of the calibration certificate? | A | B | C | D | N/A |
| How satisfied are you with the technician’s response? | A | B | C | D | N/A |
| How satisfied are you with the turn-around time for service/calibration? | A | B | C | D | N/A |
| Was the returned product packaged correctly? Please note any damage in question 5. | A | B | C | D | N/A |

- How likely are you to recommend Troxler to others?  (On a scale of 0-9 with 0 being never and 9 being very likely)  0       1       2       3       4       5       6       7       8       9

- Do you require that calibration providers be ISO/IEC 17025 accredited?    Yes      No

- What other calibration services would you like us to offer?

| Calipers | Other (please list): |  |
| --- | --- | --- |
| Compression machines |  |  |
| Dial gauges |  |  |
| Scales & weights |  |  |
| Thermometers |  |  |

- How can we better serve your calibration needs?

- How may we contact you?

Name: _______________________________________   Company: __________________________________

Phone: ______________________________________   Email: _____________________________________

| Thank you for responding to this survey. Your input is important to us as we work to improve our quality, customer service, and measurement capabilities. Please fax completed form to (919) 549-0761, or scan and email to quality@troxlerlabs.com, or submit online at the QRC to the right. |  |
| --- | --- |