## 76636
3450 Nuclear Density Gauges (Intial Quality Check)
Date: | 12/10/2019
Tech ID: | SPM |  |  | S/N: | 76636
Stat Tests |  |  |  | Projected Counts (Projected Decay Sheet)
SYS 1 | SYS 2 | Moisture |  | Month | SYS 1 | SYS2 | Moisture
6676 | 2021 | 1335 |  | 12/2019 | 6668 | 2013 | 1333
6674 | 2014 | 1333
6676 | 2018 | 1338 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3
SYS1 | SYS2 | Moisture
6679 | 2012 | 1337
6677 | 2014 | 1335
6670 | 2019 | 1338
Standard Counts w/ Difference From Projected
SYS 1 | Pass/Fail | % Diff | SYS 2 | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
6682 | P | 0.20996 | 2028 | P | 0.74516 | 1347 | F | 1.0503
6665 | P | -0.044991 | 2016 | P | 0.14903 | 1329 | P | -0.30008
6647 | P | -0.31494 | 2005 | P | -0.39742 | 1342 | P | 0.67517
Source Rod Quickly Indexed (8" - Safe) X3
6694 | P | 0.38992 | 1990 | P | -1.1426 | 1329 | P | -0.30008
6668 | P | 0 | 2008 | P | -0.24839 | 1337 | P | 0.30008
6666 | P | -0.029994 | 2025 | P | 0.59613 | 1335 | P | 0.15004
Quality Check Measurements
|  | Mag Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF) | 4 Min Thin Layer Density Counts
[9 empty] | 1.00" | 2.00"
1 Min | B/S | 109.6 |  | 164.2 |  | 35.5 | Magnesium (106.7-113.3) |  | 110.3 | 107.7
15 Sec | 2" | 109.4 |  | 162.2 |  |  | Mg / Al (134.2-139.0) |  | 134.3 | 135.0
| 4" | 109.8 |  | 163.6 |  |  | Aluminium (158.6-162.8) |  | 158.8 | 159.4
| 6" | 109.8 |  | 163.5
| 8" | 109.5 |  | 163.3

## 76637
3450 Nuclear Density Gauges (Intial Quality Check)
Date: | 12/11/2019
Tech ID: | SPM |  |  | S/N: | 76637
Stat Tests |  |  |  | Projected Counts (Projected Decay Sheet)
SYS 1 | SYS 2 | Moisture |  | Month | SYS 1 | SYS2 | Moisture
3732 | 2079 | 1263 |  | 12/2019 | 6638 | 2074 | 1251
3757 | 2091 | 1257
3753 | 2093 | 1262 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  |  | SYS 1 consistently counting 44% low. Sending back
SYS1 | SYS2 | Moisture |  | to Troxler.
3719 | 2078 | 1262
Standard Counts w/ Difference From Projected
SYS 1 | Pass/Fail | % Diff | SYS 2 | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
| P | -100 |  | P | -100 |  | P | -100
| P | -100 |  | P | -100 |  | P | -100
| P | -100 |  | P | -100 |  | P | -100
Source Rod Quickly Indexed (8" - Safe) X3
| P | -100 |  | P | -100 |  | P | -100
| P | -100 |  | P | -100 |  | P | -100
| P | -100 |  | P | -100 |  | P | -100
Quality Check Measurements
|  | Mag Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF) | 4 Min Thin Layer Density Counts
[9 empty] | 1.00" | 2.00"
1 Min | B/S | [5 empty] | Magnesium (106.7-113.3)
15 Sec | 2" | [5 empty] | Mg / Al (134.2-139.0)
| 4" | [5 empty] | Aluminium (158.6-162.8)
6"
8"

## 76638
3450 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/11/2019
Tech ID: | SPM |  |  | S/N: | 76638
Stat Tests |  |  |  | Projected Counts (Projected Decay Sheet)
SYS 1 | SYS 2 | Moisture |  | Month | SYS 1 | SYS2 | Moisture
6632 | 1931 | 1275 |  | 12/2019 | 6592 | 1928 | 1276
6633 | 1930 | 1273
6633 | 1927 | 1273 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3
SYS1 | SYS2 | Moisture
6639 | 1935 | 1276
6637 | 1933 | 1273
6639 | 1935 | 1275
Standard Counts w/ Difference From Projected
SYS 1 | Pass/Fail | % Diff | SYS 2 | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
6634 | P | 0.63714 | 1925 | P | -0.1556 | 1277 | P | 0.07837
6626 | P | 0.51578 | 1943 | P | 0.77801 | 1280 | P | 0.31348
6634 | P | 0.63714 | 1933 | P | 0.25934 | 1266 | P | -0.7837
Source Rod Quickly Indexed (8" - Safe) X3
6649 | P | 0.86468 | 1944 | P | 0.82988 | 1275 | P | -0.07837
6633 | P | 0.62197 | 1927 | P | -0.051867 | 1288 | P | 0.94044
6630 | P | 0.57646 | 1932 | P | 0.20747 | 1282 | P | 0.47022
Quality Check Measurements
|  | Mag Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF) | 4 Min Thin Layer Density Counts
[9 empty] | 1.00" | 2.00"
1 Min | B/S | 111.1 |  | 163.3 |  | 36.8 | Magnesium (106.7-113.3) |  | 112.7 | 111.6
15 Sec | 2" | 109.8 |  | 163.4 |  |  | Mg / Al (134.2-139.0) |  | 138.9 | 137.3
| 4" | 109.7 |  | 164.1 |  |  | Aluminium (158.6-162.8) |  | 161.0 | 161.9
| 6" | 110 |  | 163.8
| 8" | 109.5 |  | 162.6

## 76639
3450 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/12/2019
Tech ID: | SPM |  |  | S/N: | 76639
Stat Tests |  |  |  | Projected Counts (Projected Decay Sheet)
SYS 1 | SYS 2 | Moisture |  | Month | SYS 1 | SYS2 | Moisture
7561 | 2039 | 1178 |  | 12/2019 | 7473 | 2017 | 1180
7557 | 2044 | 1179
7548 | 2044 | 1177 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  |  | SYS 1 & SYS 2 scewed high and Std Cnts are outside of
SYS1 | SYS2 | Moisture |  | calibration projected decay. Gauge reads high for
7572 | 2044 | 1176 |  | Moisture Quality but barely within tolerance. Mag
7549 | 2044 | 1177 |  | TL 1.00" Quality reads high. 2.00" reads within tolerance.
7554 | 2044 | 1175
Standard Counts w/ Difference From Projected
SYS 1 | Pass/Fail | % Diff | SYS 2 | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
7566 | P | 1.2445 | 2035 | P | 0.89241 | 1195 | F | 1.2712
7594 | F | 1.6192 | 2049 | F | 1.5865 | 1175 | P | -0.42373
7561 | P | 1.1776 | 2064 | F | 2.3302 | 1182 | P | 0.16949
Source Rod Quickly Indexed (8" - Safe) X3
7556 | P | 1.1107 | 2028 | P | 0.54536 | 1171 | P | -0.76271
7598 | F | 1.6727 | 2025 | P | 0.39663 | 1184 | P | 0.33898
7560 | P | 1.1642 | 2024 | P | 0.34705 | 1180 | P | 0
Quality Check Measurements
|  | Mag Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF) | 4 Min Thin Layer Density Counts
[9 empty] | 1.00" | 2.00"
1 Min | B/S | 110.2 |  | 163.4 |  | 37.0 | Magnesium (106.7-113.3) |  | 117.1 | 113.2
15 Sec | 2" | 110.4 |  | 163 |  |  | Mg / Al (134.2-139.0) |  | 138.7 | 138.8
| 4" | 110.1 |  | 163.6 |  |  | Aluminium (158.6-162.8) |  | 162.3 | 162.0
| 6" | 110.8 |  | 162.6
| 8" | 109.9 |  | 163.4

## 76640
3450 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/12/2019
Tech ID: | SPM |  |  | S/N: | 76640
Stat Tests |  |  |  | Projected Counts (Projected Decay Sheet)
SYS 1 | SYS 2 | Moisture |  | Month | SYS 1 | SYS2 | Moisture
6767 | 1960 | 1299 |  | 12/2019 | 6753 | 1976 | 1303
6761 | 1960 | 1303
6757 | 1962 | 1304 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  |  | Alk Batts Sys 1 & 2 reading low w/ new batts installed.
SYS1 | SYS2 | Moisture |  | Alk Sys 2 reads 0.0 V. Std Cnts fluctuating a bit. Alum
6745 | 1960 | 1297 |  | Quality Check counts high @ 4" & 6". Took many recounts
6742 | 1962 | 1307 |  | at 1 minute. Mag TL @ 1.00" & 2.00"high. Results consistent.
6749 | 1962 | 1297
Standard Counts w/ Difference From Projected
SYS 1 | Pass/Fail | % Diff | SYS 2 | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
6737 | P | -0.23693 | 1969 | P | -0.35425 | 1299 | P | -0.30698
6751 | P | -0.029616 | 1973 | P | -0.15182 | 1285 | F | -1.3814
6681 | P | -1.0662 | 1949 | F | -1.3664 | 1298 | P | -0.38373
Source Rod Quickly Indexed (8" - Safe) X3
6751 | P | -0.029616 | 1968 | P | -0.40486 | 1307 | P | 0.30698
6688 | P | -0.96254 | 1942 | F | -1.7206 | 1304 | P | 0.076746
6738 | P | -0.22212 | 1966 | P | -0.50607 | 1300 | P | -0.23024
Quality Check Measurements
|  | Mag Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF) | 4 Min Thin Layer Density Counts
[9 empty] | 1.00" | 2.00"
1 Min | B/S | 110.1 |  | 162.6 |  | 36.6 | Magnesium (106.7-113.3) |  | 113.7 | 113.7
15 Sec | 2" | 110.2 |  | 162.4 |  |  | Mg / Al (134.2-139.0) |  | 138.0 | 136.0
| 4" | 110.5 |  | 165.1 |  |  | Aluminium (158.6-162.8) |  | 161.5 | 160.7
| 6" | 109.9 |  | 167.5
| 8" | 109.9 |  | 162.6

## 76641
3450 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/13/2019
Tech ID: | SPM |  |  | S/N: | 76641
Stat Tests |  |  |  | Projected Counts (Projected Decay Sheet)
SYS 1 | SYS 2 | Moisture |  | Month | SYS 1 | SYS2 | Moisture
7307 | 2002 | 1215 |  | 12/2019 | 7288 | 2006 | 1212
7306 | 2000 | 1216
7312 | 1994 | 1220 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  |  | Std Cnts drifting in and out of tolerance. Density Systems
SYS1 | SYS2 | Moisture |  | dropped after quicklu indexing source rod. Bioshield bushing?
7232 | 1987 | 1220 |  | Calibration is questionable. Many recounts out of tolerance.
7308 | 1999 | 1217
7310 | 2003 | 1218
Standard Counts w/ Difference From Projected
SYS 1 | Pass/Fail | % Diff | SYS 2 | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
7293 | P | 0.068606 | 2007 | P | 0.04985 | 1229 | F | 1.4026
7300 | P | 0.16465 | 1994 | P | -0.59821 | 1216 | P | 0.33003
7258 | P | -0.41164 | 1967 | F | -1.9442 | 1222 | P | 0.82508
Source Rod Quickly Indexed (8" - Safe) X3
7176 | F | -1.5368 | 1979 | F | -1.346 | 1211 | P | -0.082508
7186 | F | -1.3996 | 1983 | P | -1.1466 | 1218 | P | 0.49505
7199 | P | -1.2212 | 1979 | F | -1.346 | 1227 | F | 1.2376
Quality Check Measurements
|  | Mag Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF) | 4 Min Thin Layer Density Counts
[9 empty] | 1.00" | 2.00"
1 Min / 4 Min RC | B/S | 109.7 |  | 160.6 |  | 35.1 | Magnesium (106.7-113.3) |  | 107.5 | 108.9
15 Sec / 1 Min RC | 2" | 109.3 |  | 164 |  |  | Mg / Al (134.2-139.0) |  | 130.8 | 130.2
| 4" | 109.5 |  | 165.9 |  |  | Aluminium (158.6-162.8) |  | 158.9 | 156.9
| 6" | 108.8 |  | 162.2
| 8" | 109.6 |  | 163

## 76642
3450 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/16/2019
Tech ID: | SPM |  |  | S/N: | 76642
Stat Tests |  |  |  | Projected Counts (Projected Decay Sheet)
SYS 1 | SYS 2 | Moisture |  | Month | SYS 1 | SYS2 | Moisture
7160 | 2050 | 1258 |  | 12/2019 | 7094 | 2026 | 1247
7164 | 2046 | 1259
7169 | 2049 | 1256 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  |  | Alk Batts Sys 1 & 2 reading low w/ new batts installed.
SYS1 | SYS2 | Moisture |  | Alk Sys 2 reads 0.0 V. Depth strip will not index @ BS for
7166 | 2040 | 1261 |  | auto depth calibration. Al Density QA well out of tolerance
7169 | 2045 | 1253 |  | for 6" & 8".
7161 | 2044 | 1258
Standard Counts w/ Difference From Projected
SYS 1 | Pass/Fail | % Diff | SYS 2 | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
7164 | P | 0.98675 | 2055 | F | 1.4314 | 1260 | P | 1.0425
7107 | P | 0.18325 | 2021 | P | -0.24679 | 1258 | P | 0.88212
7139 | P | 0.63434 | 2043 | P | 0.83909 | 1254 | P | 0.56135
Source Rod Quickly Indexed (8" - Safe) X3
7136 | P | 0.59205 | 2041 | P | 0.74038 | 1257 | P | 0.80192
7147 | P | 0.74711 | 2038 | P | 0.5923 | 1256 | P | 0.72173
7116 | P | 0.31012 | 2045 | P | 0.93781 | 1261 | F | 1.1227
Quality Check Measurements
|  | Mag Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF) | 4 Min Thin Layer Density Counts
[9 empty] | 1.00" | 2.00"
1 Min | B/S | 110.1 |  | 164.2 |  | 35.0 | Magnesium (106.7-113.3) |  | 109.7 | 108.7
15 Sec | 2" | 110.6 |  | 163.2 |  |  | Mg / Al (134.2-139.0) |  | 134.9 | 134.7
| 4" | 110.4 |  | 164.2 |  |  | Aluminium (158.6-162.8) |  | 161.0 | 160.7
| 6" | 110.3 |  | 168.6
| 8" | 110 |  | 167.2

## 76643
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/16/2019
Tech ID: | SPM |  |  | S/N: | 76643
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2634 | 713 |  | 12/2019 | 2616 | 706
2628 | 713
2639 | 710 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | Density Std Cnt drifted over 1% after indexing source rod X3
Density | Moisture |  | Bioshield bushing issue? Still reads accurately.
2616 | 714
2627 | 710
2617 | 711
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2625 | P | 0.34404 | 712 | P | 0.84986
2625 | P | 0.34404 | 715 | P | 1.2748
2619 | P | 0.11468 | 712 | P | 0.84986
Source Rod Quickly Indexed (8" - Safe) X3
2651 | F | 1.3379 | 715 | P | 1.2748
2650 | F | 1.2997 | 715 | P | 1.2748
2657 | F | 1.5673 | 713 | P | 0.9915
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 110.5 |  | 162.5 |  | 36.2
15 Sec | 2" | 110 |  | 162.5
| 4" | 109.9 |  | 163.1
| 6" | 109.8 |  | 162.8
| 8" | 109.5 |  | 162.9

## 76644
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/17/2019
Tech ID: | SPM |  |  | S/N: | 76644
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2782 | 736 |  | 12/2019 | 2748 | 734
2783 | 734
2788 | 737 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | Density Std Cnt drifted high after source rod indexed. Quality check
Density | Moisture |  | still accurate.
2765 | 736
2756 | 736
2760 | 733
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2754 | P | 0.21834 | 741 | P | 0.95368
2757 | P | 0.32751 | 733 | P | -0.13624
2760 | P | 0.43668 | 737 | P | 0.40872
Source Rod Quickly Indexed (8" - Safe) X3
2770 | P | 0.80058 | 730 | P | -0.54496
2779 | F | 1.1281 | 736 | P | 0.27248
2778 | F | 1.0917 | 737 | P | 0.40872
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 111 |  | 163.4 |  | 36.5
15 Sec | 2" | 109.8 |  | 162.7
| 4" | 109.2 |  | 163.8
| 6" | 111.1 |  | 163.5
| 8" | 109.5 |  | 162.5

## 76645
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/17/2019
Tech ID: | SPM |  |  | S/N: | 76645
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2570 | 775 |  | 12/2019 | 2522 | 772
2576 | 771
2580 | 773 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | Density Std was out of tolerance prior to cycling index rod.
Density | Moisture |  | Mag BS low. Many recounts.
2551 | 775
2549 | 776
2555 | 774
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2559 | F | 1.4671 | 771 | P | -0.12953
2555 | F | 1.3085 | 785 | P | 1.6839
2555 | F | 1.3085 | 773 | P | 0.12953
Source Rod Quickly Indexed (8" - Safe) X3
2514 | P | -0.31721 | 775 | P | 0.3886
2511 | P | -0.43616 | 768 | P | -0.51813
2515 | P | -0.27756 | 771 | P | -0.12953
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 108.9 |  | 162.5 |  | 36.2
15 Sec | 2" | 109.2 |  | 162.2
| 4" | 109.5 |  | 162.6
| 6" | 109.3 |  | 163
| 8" | 109.5 |  | 163.2

## 76646
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/18/2019
Tech ID: | SPM |  |  | S/N: | 76646
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2856 | 791 |  | 12/2019 | 2816 | 785
2860 | 792
2858 | 791 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | Density Std Cnts ran high. Gauge reads accurately.
Density | Moisture
2844 | 793
2839 | 795
2841 | 793
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2836 | P | 0.71023 | 794 | P | 1.1465
2844 | P | 0.99432 | 793 | P | 1.0191
2856 | F | 1.4205 | 790 | P | 0.63694
Source Rod Quickly Indexed (8" - Safe) X3
2859 | F | 1.527 | 785 | P | 0
2843 | P | 0.95881 | 796 | P | 1.4013
2857 | F | 1.456 | 791 | P | 0.76433
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 110.4 |  | 162.7 |  | 35.8
15 Sec | 2" | 109.9 |  | 163
| 4" | 109.5 |  | 163.9
| 6" | 110.9 |  | 163.8
| 8" | 110.2 |  | 164

## 76647
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/18/2019
Tech ID: | SPM |  |  | S/N: | 76647
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2696 | 701 |  | 12/2019 | 2664 | 694
2694 | 699
2696 | 700 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | Density Std Cnt drifted a bit. Density Quality Check for Mag and Alum
Density | Moisture |  | 2" and 4" reading low density.
2665 | 696
2660 | 699
2657 | 700
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2665 | P | 0.037538 | 696 | P | 0.28818
2638 | P | -0.97598 | 691 | P | -0.43228
2644 | P | -0.75075 | 701 | P | 1.0086
Source Rod Quickly Indexed (8" - Safe) X3
2634 | F | -1.1261 | 702 | P | 1.1527
2635 | F | -1.0886 | 695 | P | 0.14409
2645 | P | -0.71321 | 697 | P | 0.43228
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 109.5 |  | 162.8 |  | 35.3
15 Sec | 2" | 108.5 |  | 161.6
| 4" | 109.1 |  | 161.7
| 6" | 109.4 |  | 162.2
| 8" | 109.3 |  | 162.5

## 76648
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/19/2019
Tech ID: | SPM |  |  | S/N: | 76648
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2468 | 715 |  | 12/2019 | 2495 | 711
2467 | 714
2464 | 715 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3
Density | Moisture
2490 | 715
2484 | 714
2487 | 717
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2485 | P | -0.4008 | 723 | P | 1.6878
2495 | P | 0 | 720 | P | 1.2658
2483 | P | -0.48096 | 711 | P | 0
Source Rod Quickly Indexed (8" - Safe) X3
2501 | P | 0.24048 | 711 | P | 0
2503 | P | 0.32064 | 715 | P | 0.56259
2493 | P | -0.08016 | 712 | P | 0.14065
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 109.2 |  | 163 |  | 35.3
15 Sec | 2" | 109.4 |  | 162.6
| 4" | 109.2 |  | 162.5
| 6" | 109.2 |  | 162.2
| 8" | 109.2 |  | 163.4

## 76649
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/19/2019
Tech ID: | SPM |  |  | S/N: | 76649
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2808 | 735 |  | 12/2019 | 2747 | 736
2802 | 735
2804 | 734 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | Density Std drifted over night. Alum BS Quality Check was low.
Density | Moisture
2775 | 735
2773 | 734
2777 | 735
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2768 | P | 0.76447 | 733 | P | -0.40761
2765 | P | 0.65526 | 733 | P | -0.40761
2769 | P | 0.80087 | 731 | P | -0.67935
Source Rod Quickly Indexed (8" - Safe) X3
2758 | P | 0.40044 | 724 | P | -1.6304
2767 | P | 0.72807 | 735 | P | -0.13587
2789 | F | 1.5289 | 739 | P | 0.40761
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 110 |  | 160.7 |  | 35.9
15 Sec | 2" | 110 |  | 164.2
| 4" | 109.4 |  | 163.2
| 6" | 110.6 |  | 162.6
| 8" | 109.7 |  | 163

## 76650
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/20/2019
Tech ID: | SPM |  |  | S/N: | 76650
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2795 | 724 |  | 12/2019 | 2724 | 721
2803 | 723
2798 | 725 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | Density Std Cnt out of tolerance, high for projected decay.
Density | Moisture |  | All Quality Check densities in tolerance.
2762 | 723
2762 | 724
2766 | 725
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2765 | F | 1.5051 | 717 | P | -0.55479
2754 | F | 1.1013 | 725 | P | 0.55479
2770 | F | 1.6887 | 731 | P | 1.387
Source Rod Quickly Indexed (8" - Safe) X3
2762 | F | 1.395 | 722 | P | 0.1387
2755 | F | 1.138 | 725 | P | 0.55479
2757 | F | 1.2115 | 725 | P | 0.55479
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 110.9 |  | 163.2 |  | 36.6
15 Sec | 2" | 110.6 |  | 162.4
| 4" | 110.3 |  | 163
| 6" | 110.1 |  | 163.7
| 8" | 109.4 |  | 163.4

## 76651
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/20/2019
Tech ID: | SPM |  |  | S/N: | 76651
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
2881 | 719 |  | 12/2019 | 2822 | 718
2888 | 723
2882 | 723 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | Most density Std Cnts beyond decay tolerance. Gauge reads accurate
Density | Moisture |  | however.
2860 | 720
2862 | 721
2860 | 719
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
2849 | F | 0.95677 | 727 | P | 1.2535
2864 | F | 1.4883 | 713 | P | -0.69638
2845 | P | 0.81502 | 721 | P | 0.41783
Source Rod Quickly Indexed (8" - Safe) X3
2881 | F | 2.0907 | 723 | P | 0.69638
2877 | F | 1.949 | 723 | P | 0.69638
2879 | F | 2.0198 | 725 | P | 0.97493
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 110.4 |  | 164.1 |  | 35.5
15 Sec | 2" | 109.7 |  | 164.2
| 4" | 110.1 |  | 162.6
| 6" | 110.1 |  | 162.4
| 8" | 110 |  | 163.6

## 76652
3430 Nuclear Density Gauges (Initial Quality Check)
Date: | 12/27/2019
Tech ID: | SPM |  |  | S/N: | 76652
Stat Tests |  |  | Projected Counts (Projected Decay Sheet)
Density | Moisture |  | Month | Density | Moisture
3064 | 741 |  | 12/2019 | 3050 | 734
3062 | 740
3064 | 742 |  | Notes:
Source Rod Quickly Indexed (8" - Safe) X3 |  |  | A few density Std Cnts out of tolerance. Gauge reads accurately
Density | Moisture |  | except @ Alum BS.
3072 | 738
3070 | 740
3072 | 745
Standard Counts w/ Difference From Projected
Density | Pass/Fail | % Diff | Moisture | Pass/Fail | % Diff
3072 | P | 0.72131 | 741 | P | 0.95368
3070 | P | 0.65574 | 740 | P | 0.81744
3074 | P | 0.78689 | 743 | P | 1.2262
Source Rod Quickly Indexed (8" - Safe) X3
3098 | F | 1.5738 | 744 | P | 1.3624
3090 | F | 1.3115 | 741 | P | 0.95368
3058 | P | 0.2623 | 736 | P | 0.27248
Quality Check Measurements
|  | Magnesium Density (109.2-111.2 PCF) |  | Aluminum Density (162.2-164.2 PCF) |  | Moisture (35.0-37.0 PCF)
1 Min | B/S | 110.6 |  | 166 |  | 36.1
15 Sec | 2" | 109.4 |  | 164.2
| 4" | 109.8 |  | 164.2
| 6" | 110.2 |  | 162.8
| 8" | 110.5 |  | 164